FPGA-Based Diabetic Retinopathy Detection System

Author

Salamat Ali

Muhammad Shahroz Chowdry

A thesis submitted in partial fulfillment of the requirements for the degree of B.Sc. Electronic Engineering

Supervisor

Dr. Dileep Kumar

Assistant Professor

Session: 2021-2025

External Examiner Signature:	______________________

Thesis Supervisor Signature:	______________________

DEPARTMENT OF ELECTRONIC ENGINEERING

UNIVERSITY COLLEGE OF ENGINEERING AND TECHNOLOGY

The Islamia University of Bahawalpur

# Abstract

Diabetic Retinopathy (DR) is one of the leading causes of preventable blindness among working-age adults worldwide, with an estimated 93 million people currently affected globally. Early and accurate detection remains the most effective intervention, yet the scarcity of ophthalmologists in rural and resource-limited settings—particularly in countries such as Pakistan—presents a formidable barrier to timely screening. This thesis presents a low-cost, low-power, FPGA-based automated DR detection system designed to bridge this critical healthcare gap.

The proposed system employs a lightweight Convolutional Neural Network (CNN) trained on retinal fundus images to classify DR severity into five grades (0–4) according to the standard clinical scale. The model is quantized to INT8 precision and compiled for deployment on a custom Deep Processing Unit (DPU) integrated within the Xilinx PYNQ-Z2 heterogeneous System-on-Chip (SoC), which combines a dual-core ARM Cortex-A9 processor with Artix-7 FPGA fabric. A cloud-based pipeline built on Firebase enables remote image upload via a Streamlit web interface, facilitating telemedicine-ready screening without requiring physical access to specialist equipment.

The system achieves an FPGA inference latency of 15–25 milliseconds per image and an end-to-end pipeline latency of 2–3 seconds, while consuming only 6–8 Watts of power—enabling operation from battery or solar sources in off-grid environments. The model attains a validation accuracy of 87–92% on a benchmark retinal image dataset. FPGA resource utilization remains within practical bounds: 70–75% BRAM, 60–65% DSP slices, and approximately 71% LUT occupancy. The total hardware cost is approximately $200, compared to $10,000–$50,000 for conventional fundus camera and workstation setups.

This work establishes a reproducible methodology for deploying medical-grade deep learning inference on low-cost FPGA hardware, and demonstrates the viability of edge AI for primary healthcare applications in developing regions. The integration of cloud storage, real-time OLED output, and a web-accessible interface provides a complete, clinically deployable screening solution.

Keywords: Diabetic Retinopathy; Edge Artificial Intelligence; FPGA Acceleration; Convolutional Neural Network; Deep Processing Unit (DPU); INT8 Quantization; PYNQ-Z2; Firebase Cloud Integration; Telemedicine; Low-Power Embedded Systems; Healthcare IoT.

# Undertaking

We certify that the research work titled "FPGA-Based Diabetic Retinopathy Detection System" is our original work. The work has not been presented elsewhere for assessment. Where material has been used from other sources, it has been properly acknowledged and referenced.

Salamat Ali

______________________

Muhammad Shahroz Chowdry

______________________

# Acknowledgements

We would like to sincerely thank our supervisor, Dr. Dileep Kumar, for his invaluable guidance, unwavering patience, and continuous encouragement throughout the duration of this research. His deep expertise in deep learning, FPGA design, and embedded systems significantly shaped the direction and quality of this work.

We are also grateful to the Department of Electronic Engineering, University College of Engineering and Technology, The Islamia University of Bahawalpur, for providing the laboratory facilities, computational resources, and institutional support that made this project possible. The assistance of the technical staff in the FPGA laboratory—particularly with board configuration, PetaLinux integration, and networking setup—was instrumental in overcoming several hardware challenges.

We extend our gratitude to the Xilinx community and the open-source contributors whose documentation, example designs, and forum discussions provided essential guidance on DPU configuration and Vitis AI toolchain usage.

Additionally, we would like to acknowledge the creators and maintainers of the APTOS 2019 Blindness Detection and EyePACS datasets, whose publicly available retinal fundus image collections formed the foundation of our model training pipeline.

Finally, we are profoundly thankful to our families for their constant encouragement, understanding, and belief in our abilities throughout this academic journey. Their support has been the most enduring motivation behind this work.

# Table of Contents

# List of Figures

Figure 1.1:  High-level block diagram of the proposed DR detection system.	XX

Figure 1.2:  End-to-end data flow from image upload to FPGA inference and result display.	XX

Figure 2.1:  Architecture of the Xilinx Zynq-7000 SoC.	XX

Figure 3.1:  PYNQ-Z2 board physical layout and peripheral connectivity.	XX

Figure 3.2:  PS-PL communication via AXI interconnect.	XX

Figure 3.3:  DPU block design within Vivado IP Integrator.	XX

Figure 3.4:  Configured DPU final design in Vivado.	XX

Figure 3.5:  Custom OS build flow in PetaLinux.	XX

Figure 3.6:  Lightweight CNN architecture for DR classification.	XX

Figure 3.7:  Depthwise separable convolution block diagram.	XX

Figure 3.8:  INT8 quantization workflow using Vitis AI DECENT_Q.	XX

Figure 3.9:  Firebase cloud architecture showing Storage, Firestore, and Auth.	XX

Figure 3.10:  Streamlit web interface for retinal image upload.	XX

Figure 3.11:  OLED display showing real-time prediction result.	XX

Figure 3.12:  System connectivity and network topology.	XX

Figure 4.1:  Model training loss and accuracy curves.	XX

Figure 4.2:  Confusion matrix for 5-class DR classification.	XX

Figure 4.3:  Per-class precision-recall curves.	XX

Figure 4.4:  Latency comparison: FPGA vs CPU inference.	XX

Figure 4.5:  Power consumption during idle and active inference.	XX

Figure 4.6:  Sample detection results on retinal fundus images.	XX

# List of Tables

Table 1.1:  Comparison of screening cost: traditional vs proposed system.	XX

Table 1.2:  Summary of scope and limitations.	XX

Table 2.1:  Benchmark summary of related FPGA-based CNN deployments.	XX

Table 3.1:  PYNQ-Z2 board hardware specifications.	XX

Table 3.2:  DPU resource utilization after implementation.	XX

Table 3.3:  CNN model architecture with layer parameters.	XX

Table 3.4:  Training configuration hyperparameters.	XX

Table 3.5:  Firebase Firestore collection schema.	XX

Table 4.1:  Classification accuracy per DR grade.	XX

Table 4.2:  Inference performance comparison: FPGA vs CPU.	XX

Table 4.3:  Power consumption budget breakdown.	XX

# Abbreviations

AXI:  Advanced Extensible Interface

BRAM:  Block Random Access Memory

CNN:  Convolutional Neural Network

CPU:  Central Processing Unit

DECENT_Q:  Deep Compression and Elimination Tool (Quantization)

DPU:  Deep Processing Unit

DR:  Diabetic Retinopathy

DSP:  Digital Signal Processing

FPGA:  Field Programmable Gate Array

GPIO:  General Purpose Input/Output

GPU:  Graphical Processing Unit

HDMI:  High-Definition Multimedia Interface

I2C:  Inter-Integrated Circuit

IDF:  International Diabetes Federation

INT8:  8-bit Integer Quantization

IoT:  Internet of Things

IP:  Intellectual Property (block)

NPDR:  Non-Proliferative Diabetic Retinopathy

OLED:  Organic Light-Emitting Diode

PL:  Programmable Logic

PS:  Processing System

PYNQ:  Python Productivity for Zynq

ReLU:  Rectified Linear Unit

REST:  Representational State Transfer

SoC:  System-on-Chip

SSD:  Single Shot MultiBox Detector

USB:  Universal Serial Bus

XSA:  Xilinx Shell Archive

# CHAPTER 1: INTRODUCTION

## 1.1 Background

Diabetes mellitus is one of the most prevalent chronic diseases of the modern era. According to the International Diabetes Federation (IDF), approximately 537 million adults were living with diabetes in 2021, a figure projected to reach 783 million by 2045. Among the many complications of prolonged diabetes, Diabetic Retinopathy (DR) represents one of the most clinically significant, being a leading cause of acquired blindness in working-age adults. DR is caused by progressive damage to the microvasculature of the retina resulting from chronically elevated blood glucose levels, and affects an estimated 93 million people globally.

The clinical progression of DR is well-characterized and spans five distinct stages, from no apparent retinopathy (Grade 0) through mild, moderate, and severe non-proliferative stages (Grades 1–3), to proliferative DR (Grade 4), in which pathological neovascularization and vitreous hemorrhage can lead to irreversible vision loss. Crucially, vision loss from DR is largely preventable: regular screening to detect early-stage disease, followed by timely intervention, can prevent approximately 95% of severe vision loss cases. However, the effectiveness of screening programs depends heavily on the availability of trained ophthalmologists and diagnostic infrastructure—resources that remain severely limited in low- and middle-income countries.

In Pakistan specifically, the ratio of ophthalmologists to population in rural areas is estimated at 1 per 100,000, far below the World Health Organization's recommended threshold. The vast majority of specialist eye-care services are concentrated in urban centers, leaving rural diabetic patients with little or no access to retinal screening. This disparity results in a disproportionate burden of late-stage DR diagnoses, where the window for effective treatment has often already closed.

Advances in artificial intelligence—particularly deep learning-based computer vision—have demonstrated the potential to automate DR screening from digital fundus photographs with accuracy comparable to trained specialists. However, most such systems are designed for deployment on powerful GPU-based servers with continuous cloud connectivity, making them unsuitable for resource-constrained clinical settings. There is a compelling need for automated DR detection systems that are simultaneously accurate, affordable, power-efficient, and operable in environments with limited infrastructure.

Field-Programmable Gate Arrays (FPGAs) have emerged as a promising hardware platform for edge AI applications, offering a favorable balance of computational throughput, energy efficiency, and reconfigurability compared to both CPUs and dedicated AI accelerators. The Xilinx PYNQ-Z2 is a low-cost FPGA development board built around the Zynq-7000 SoC, which integrates a dual-core ARM Cortex-A9 processor (Processing System, PS) with Artix-7 FPGA fabric (Programmable Logic, PL). This heterogeneous architecture is well-suited for the hardware-software co-design required to deploy quantized deep learning inference at the edge.

[IMAGE PLACEHOLDER] Figure 1.1: High-level block diagram of the proposed FPGA-based DR detection system.

Figure 1.1: High-level block diagram of the proposed FPGA-based DR detection system.

## 1.2 Problem Statement

Despite the clinical efficacy of automated DR detection using deep learning, several technical and logistical barriers prevent the widespread adoption of such systems in rural healthcare settings. The core challenges may be summarized as follows:

- High computational cost and hardware dependency: State-of-the-art DR detection models, such as those based on ResNet-50 or EfficientNet, require GPU-class hardware for real-time inference, with associated costs of $2,000–$10,000 per workstation and power consumption exceeding 200W.

- Internet dependency: Cloud-based inference solutions require stable broadband connectivity, which is frequently unavailable in rural primary healthcare clinics in Pakistan and similar settings.

- Cost of screening equipment: Commercial fundus cameras capable of capturing the retinal images required for DR diagnosis typically cost $10,000–$50,000, and are rarely available outside tertiary care centers.

- Operational complexity: Existing automated systems require technically trained operators for setup, calibration, and maintenance, limiting their deployment in settings staffed primarily by general practitioners or paramedics.

- Lack of longitudinal patient tracking: Point-of-care screening without integrated record-keeping prevents the longitudinal monitoring necessary for effective disease management.

This thesis addresses these barriers through the design and implementation of a complete FPGA-based DR detection pipeline that is low-cost, low-power, locally inferenced, and remotely accessible via a web interface.

| Problem | Impact | Proposed Solution |
| --- | --- | --- |
| High hardware cost | Unaffordable for rural clinics | $200 PYNQ-Z2 board |
| Internet dependence | Fails in remote areas | Local FPGA inference |
| High power consumption | Grid dependency | 6–8W (battery/solar capable) |
| Specialist shortage | No on-site diagnosis | Automated 5-grade classification |
| Complex operation | Requires trained technician | Simple web interface |

Table 1.1: Problem-solution mapping for the proposed DR detection system.

## 1.3 Objectives

The primary and secondary objectives of this project are defined as follows:

### 1.3.1 Primary Objectives

- Design and train a lightweight CNN model capable of classifying DR severity into five grades with a validation accuracy exceeding 85%.

- Implement and configure a Deep Processing Unit (DPU) on the PYNQ-Z2 FPGA to enable hardware-accelerated inference with a per-image latency below 100 milliseconds.

- Develop a Streamlit-based web application supporting retinal image upload from any internet-connected device, with Firebase cloud storage and Firestore database integration.

- Integrate an SSD1306 OLED display for real-time on-device visualization of classification results and confidence scores.

- Maintain total BRAM utilization below 80% and DSP utilization below 75% to ensure design closure within the PYNQ-Z2 resource envelope.

### 1.3.2 Secondary Objectives

- Enable offline inference operation to support clinics in areas with intermittent connectivity.

- Provide patient history storage and longitudinal tracking through Firestore.

- Support multi-user access and role-based authentication for doctors, patients, and clinic administrators.

- Generate clinical referral recommendations based on predicted DR grade.

## 1.4 Scope and Limitations

The scope of this project encompasses the complete hardware-software co-design pipeline from model training to edge deployment and cloud integration. The system is designed to classify DR from preprocessed retinal fundus images conforming to the standard JPEG format at a minimum resolution of 224×224 pixels.

| Aspect | In Scope | Out of Scope |
| --- | --- | --- |
| Input | Retinal fundus images (JPEG/PNG) | Live fundus camera capture, video stream |
| Classification | 5-grade DR severity (Grades 0–4) | Other retinal pathologies (glaucoma, AMD) |
| Hardware | PYNQ-Z2 FPGA, OLED, Ethernet | GPU acceleration, custom ASIC design |
| Connectivity | Ethernet + Firebase cloud sync | 4G/LTE cellular modem integration |
| Interface | Streamlit web app, OLED display | Native mobile application (iOS/Android) |
| Deployment | Single-board deployment | Distributed multi-node inference |

Table 1.2: Summary of project scope and limitations.

The primary limitation of the system is that image acquisition relies on an external fundus camera or equivalent imaging device; the PYNQ-Z2 does not include an integrated imaging sensor. Additionally, model performance is bounded by the diversity of the training dataset, and generalization to images captured under significantly different clinical conditions (e.g., varying pupil dilation, lighting artifacts) may require re-training or fine-tuning on local data. The system is not a certified medical device and is intended as a decision-support tool for use under the supervision of qualified healthcare professionals.

## 1.5 Methodology Overview

The project follows a structured hardware-software co-design methodology comprising five principal phases:

- Phase 1 – Dataset Preparation and Model Training: A lightweight depthwise separable CNN is designed, trained, and validated on publicly available retinal fundus image datasets, including APTOS 2019 and EyePACS, augmented and balanced to address class imbalance across the five DR grades.

- Phase 2 – Model Quantization: The trained model is converted to INT8 fixed-point representation using the Vitis AI DECENT_Q quantization tool, with calibration performed on a representative subset of the training data to minimize accuracy degradation.

- Phase 3 – DPU Implementation: A B1024 DPU IP core is configured and synthesized within Vivado IP Integrator for the PYNQ-Z2 target. The resulting bitstream and hardware handoff file are used to build a custom PetaLinux image.

- Phase 4 – Cloud and Web Integration: A Firebase project is configured with Storage for image hosting and Firestore for metadata and result management. A Streamlit web application provides the user-facing interface for image upload and result retrieval.

- Phase 5 – System Integration and Validation: All components are integrated on the PYNQ-Z2 board. End-to-end latency, inference accuracy, power consumption, and FPGA resource utilization are measured and compared against baseline CPU performance.

[IMAGE PLACEHOLDER] Figure 1.2: End-to-end system flow process from image capture to result display.

Figure 1.2: End-to-end system flow process from image capture to result display.

## 1.6 Thesis Organization

The remainder of this thesis is organized as follows. Chapter 2 presents a comprehensive review of the relevant literature, covering deep learning for DR detection, FPGA-based CNN acceleration, and the limitations of existing approaches. Chapter 3 describes the proposed methodology in detail, including hardware architecture, software stack, model design, quantization, DPU implementation, and cloud integration. Chapter 4 presents the experimental results and performance analysis. The thesis concludes with a summary of findings and directions for future work.

# CHAPTER 2: LITERATURE REVIEW

## 2.1 Deep Learning for Diabetic Retinopathy Detection

The application of deep learning to automated DR screening has been an active area of research since the seminal work of Gulshan et al. (2016), who demonstrated that a deep convolutional neural network trained on over 128,000 retinal images could achieve a diagnostic sensitivity of 97.5% and specificity of 93.4% on a held-out test set—performance comparable to that of board-certified ophthalmologists. This work established the clinical viability of deep learning-based fundus image analysis and catalyzed significant subsequent research.

Ting et al. (2017) extended this work to a multi-ethnic Asian population, employing an ensemble of CNN models trained across multiple clinical cohorts to achieve a sensitivity of 90.5% and specificity of 91.6% for referable DR detection. Abramoff et al. (2018) developed the IDx-DR system, which became the first FDA-cleared autonomous AI diagnostic device, achieving 87.2% sensitivity and 90.7% specificity. However, IDx-DR is a proprietary commercial system costing several thousand dollars per unit, precluding deployment in low-resource settings.

Transfer learning approaches—whereby pre-trained networks such as InceptionV3, VGG-19, or EfficientNet are fine-tuned on retinal image datasets—have become increasingly popular due to their ability to achieve high accuracy with limited labeled data. Kermany et al. (2018) demonstrated that transfer learning from ImageNet-pretrained models could achieve 92.8% accuracy on retinal OCT classification with substantially reduced training data requirements. However, these models remain computationally intensive at inference time, typically requiring GPU hardware.

More recent work has focused on model compression and lightweight architectures suitable for edge deployment. MobileNet and its variants, which use depthwise separable convolutions to reduce computational complexity, have been applied to DR grading with accuracy comparable to full-size models at a fraction of the parameter count. EfficientNet-Lite, designed specifically for edge inference, has demonstrated state-of-the-art accuracy-efficiency trade-offs on mobile benchmarks.

## 2.2 Hardware Acceleration Using FPGAs

FPGAs have attracted considerable attention as inference accelerators for deep learning workloads, owing to their reconfigurability, fine-grained parallelism, and superior energy efficiency compared to general-purpose processors. Early work by Zhang et al. (2015) demonstrated that optimized CNN inference on an FPGA could achieve significant speedups over CPU baselines through techniques including loop unrolling, tiling, and on-chip memory optimization.

The introduction of high-level synthesis (HLS) tools and vendor-provided IP cores for deep learning—most notably Xilinx's Deep Processing Unit (DPU) within the Vitis AI framework—has substantially reduced the engineering effort required to deploy CNNs on FPGA hardware. The DPU is a programmable engine optimized for the inference of quantized convolutional neural networks, supporting INT8 arithmetic and a wide range of standard layer types including convolution, batch normalization, pooling, and depthwise separable convolution.

Qiu et al. (2016) reported a 15× speedup over CPU baseline for AlexNet inference on a Zynq-7000 FPGA, consuming 9.63W of power—demonstrating the energy efficiency advantage of FPGA acceleration for medium-scale CNN workloads. Zhang et al. (2020) deployed an X-ray classification model on a PYNQ board, achieving an 18× speedup relative to Raspberry Pi CPU inference. Ghasemzadeh et al. (2019) implemented retinal OCT analysis on an Ultra96 FPGA platform, reporting a 22× speedup with under 10W power consumption.

[IMAGE PLACEHOLDER] Figure 2.1: Architecture of the Xilinx Zynq-7000 SoC showing PS-PL integration.

Figure 2.1: Architecture of the Xilinx Zynq-7000 SoC showing PS-PL integration.

## 2.3 Related Work on Automated DR Detection Systems

Several prior studies have investigated the deployment of automated DR screening tools in resource-constrained settings, though few have targeted FPGA-based edge deployment. Raumviboonsuk et al. (2019) evaluated the clinical performance of a deep learning system for DR screening in rural Thailand, demonstrating high sensitivity but relying on cloud-based GPU inference with latencies of several seconds per image. Li et al. (2020) developed a mobile-phone-based DR screening application using MobileNetV2, achieving 89% accuracy but limited by the computational constraints of smartphone processors at higher-resolution inputs.

FPGA-specific implementations for medical image analysis have been demonstrated primarily for tasks such as ECG classification, chest X-ray analysis, and ultrasound image segmentation. Direct implementation of DR grading on low-cost FPGA platforms such as the PYNQ-Z2 has not been reported in the peer-reviewed literature, representing the primary research gap addressed by this thesis.

| Study | Method | Platform | Accuracy | Limitation |
| --- | --- | --- | --- | --- |
| Gulshan et al. (2016) | Deep CNN | GPU Cluster | 97.5% | Requires GPU, high cost |
| Ting et al. (2017) | Ensemble CNN | GPU Server | 94.5% | High computational cost |
| Abramoff et al. (2018) | IDx-DR (FDA) | Proprietary HW | 87.2% | Expensive, proprietary |
| Kermany et al. (2018) | Transfer Learning | GPU | 92.8% | Requires internet/GPU |
| Zhang et al. (2020) | CNN on FPGA | PYNQ | N/A (X-ray) | Not DR-specific |
| This Work | Lightweight CNN + DPU | PYNQ-Z2 | 87–92% | Limited to fundus images |

Table 2.1: Benchmark summary of related automated retinal analysis systems.

## 2.4 Comparison of CNN Architectures for Edge Deployment

A range of CNN architectures has been evaluated in the literature for suitability on edge hardware. VGG-16 and ResNet-50, while highly accurate, have parameter counts in the range of 14–25 million, making them impractical for deployment on devices with limited BRAM and DSP resources. MobileNetV1 and V2, which employ depthwise separable convolutions to decouple channel-wise and spatial filtering, achieve competitive accuracy with 3–4 million parameters. SqueezeNet achieves further parameter reduction through the use of fire modules, at some cost to accuracy on fine-grained classification tasks.

For the five-class DR grading task, the architecture proposed in this work achieves 601,409 parameters—substantially fewer than MobileNetV2's 3.4 million—while maintaining accuracy within the clinically acceptable range. This reduction is achieved through aggressive use of depthwise separable convolutions, global average pooling in place of fully connected layers for spatial feature aggregation, and moderate dropout regularization.

## 2.5 Limitations in Current Studies

A review of the existing literature reveals several consistent limitations that motivate the contributions of this thesis. First, the majority of high-accuracy DR detection systems are designed for data-center deployment and do not address the power, cost, or connectivity constraints of edge clinical environments. Second, FPGA-based medical imaging implementations that do exist have largely targeted higher-cost platforms (Ultra96, ZCU102) beyond the budget of primary care facilities in developing countries. Third, no prior work has demonstrated a complete end-to-end pipeline integrating FPGA-based local inference with cloud-based remote access for DR screening. Fourth, the clinical workflow integration aspects—including patient record management, referral generation, and multi-user access control—have received limited attention in the engineering literature.

This thesis directly addresses the first three of these gaps and provides a foundation for future work on the fourth.

# CHAPTER 3: METHODOLOGY

## 3.1 Hardware Architecture

### 3.1.1 PYNQ-Z2 Overview

The PYNQ-Z2 is a low-cost FPGA development board manufactured by Digilent, built around the Xilinx Zynq XC7Z020-1CLG400C SoC. It is designed to support the PYNQ (Python productivity for Zynq) framework, which exposes the FPGA programmable logic to high-level Python code running on the ARM processor. This programming model significantly reduces the development cycle for hardware-accelerated embedded applications.

| Specification | Details |
| --- | --- |
| SoC | Xilinx Zynq XC7Z020-1CLG400C |
| ARM Processor | Dual-core Cortex-A9 @ 650 MHz |
| DRAM | 512 MB DDR3 @ 1066 Mb/s |
| Logic Cells | 85,000 |
| LUTs | 53,200 |
| Flip-Flops | 106,400 |
| BRAM | 4.9 Mb (140 × 36Kb blocks) |
| DSP Slices | 220 |
| Interfaces | Gigabit Ethernet, USB 2.0, HDMI, PMOD, Arduino, RPi Header |
| Operating Voltage | 12V DC, barrel jack |
| Power (active) | Approximately 7.6W |

Table 3.1: Summary of PYNQ-Z2 board hardware specifications.

[IMAGE PLACEHOLDER] Figure 3.1: PYNQ-Z2 board physical layout showing key peripheral ports.

Figure 3.1: PYNQ-Z2 board physical layout showing key peripheral ports.

### 3.1.2 Zynq-7000 SoC Architecture

The Zynq-7000 SoC consists of two primary subsystems: the Processing System (PS), which contains the dual ARM Cortex-A9 cores, L1/L2 caches, DDR3 memory controller, and a set of hardened peripherals (Ethernet, USB, SPI, I2C, UART); and the Programmable Logic (PL), which provides the configurable FPGA fabric. Communication between PS and PL is mediated by the Advanced Extensible Interface (AXI) interconnect, supporting high-bandwidth data transfer for DMA-based tensor transfers between DDR3 memory and the DPU computation engine.

For this project, the DPU occupies the PL partition and communicates with the ARM processor over AXI4-MM interfaces for weight and feature-map data, and AXI4-Lite for configuration and control registers. The ARM cores execute the Python-based inference driver, preprocessing pipeline, Firebase polling loop, and OLED display driver, while the DPU handles the computationally intensive convolution and activation operations.

[IMAGE PLACEHOLDER] Figure 3.2: PS-PL communication via AXI interconnect in the Zynq-7000 SoC.

Figure 3.2: PS-PL communication via AXI interconnect in the Zynq-7000 SoC.

### 3.1.3 Deep Processing Unit (DPU) Configuration

The DPU IP core used in this project is configured as a B1024 architecture, providing 1,024 parallel multiply-accumulate (MAC) operations per clock cycle at the configured frequency of 200 MHz, yielding a theoretical peak throughput of approximately 204.8 GOPS. The B1024 configuration was selected to balance inference performance against resource utilization within the PYNQ-Z2 budget.

| Resource | Used | Available | Utilization (%) |
| --- | --- | --- | --- |
| BRAM (36Kb) | 98 | 140 | 70% |
| DSP Slices | 132 | 220 | 60% |
| LUTs | 38,000 | 53,200 | 71% |
| Flip-Flops | 75,000 | 106,400 | 70% |

Table 3.2: FPGA resource utilization after DPU deployment.

[IMAGE PLACEHOLDER] Figure 3.3: Configured DPU block design in Vivado IP Integrator.

Figure 3.3: Configured DPU block design in Vivado IP Integrator.

## 3.2 Software Stack and Tools

### 3.2.1 Operating System: PYNQ OS / PetaLinux

The base operating system deployed on the PYNQ-Z2 is a customized Ubuntu 18.04 image provided by the PYNQ project, which includes pre-configured Jupyter notebook server, PYNQ Python libraries, and peripheral drivers. For DPU deployment, a custom PetaLinux 2019.2 build was required to incorporate the Vitis AI runtime libraries and DPU kernel module into the boot image. The PetaLinux workflow involves importing the hardware description file (XSA) exported from Vivado, configuring the rootfs to include the DNNDK/Vitis AI libraries, and building the boot artifacts (BOOT.BIN, image.ub, rootfs).

[IMAGE PLACEHOLDER] Figure 3.4: Custom OS build flow in PetaLinux for DPU-enabled PYNQ image.

Figure 3.4: Custom OS build flow in PetaLinux for DPU-enabled PYNQ image.

### 3.2.2 Vivado Design Suite

Xilinx Vivado 2019.2 was used for block design, IP integration, synthesis, implementation, and bitstream generation. The design incorporates the DPU IP, Zynq PS block, AXI interconnect, clock wizard, and reset infrastructure. The DPU is configured through the IP customization dialog with the B1024 architecture, HIGH RAM usage mode, and depthwise convolution support enabled. The finalized design is exported as an XSA file for PetaLinux integration.

### 3.2.3 Vitis AI / DNNDK Toolchain

Model quantization and compilation are performed using the Xilinx Vitis AI toolchain. The DECENT_Q quantizer converts the floating-point Keras model to an INT8 fixed-point representation through a calibration-based post-training quantization process. The DNNC compiler then converts the quantized model into a DPU instruction set binary (.xmodel), which can be loaded and executed by the DPU runtime on the PYNQ-Z2.

The quantization process requires a representative calibration dataset—in this case, 1,000 randomly sampled images from the training set—to compute per-layer activation statistics used to determine optimal quantization parameters. The compiled model is structured as a graph of DPU subgraphs and CPU subgraphs, with the latter handling operations not natively supported by the DPU such as the final softmax activation.

### 3.2.4 Python Environment and Key Libraries

All inference-time Python code is developed for Python 3.8 within an Anaconda virtual environment. Key packages employed include TensorFlow 2.x for model training and export, OpenCV 4.x for image preprocessing, NumPy for tensor manipulation, firebase-admin for cloud integration, and pynq-dpu for the PYNQ-side DPU inference API. The OLED display driver is implemented using the Adafruit SSD1306 library over I2C.

## 3.3 Implementation Steps

### 3.3.1 CNN Model Design

The proposed DR classification model is a lightweight CNN based on depthwise separable convolutions, inspired by the MobileNet family of architectures. The design prioritizes three properties: sufficient representational capacity for fine-grained 5-class retinal image classification, compatibility with INT8 quantization without excessive accuracy degradation, and resource utilization within the PYNQ-Z2 DPU budget.

The network processes 224×224×3 RGB retinal images through an initial standard 3×3 convolution with stride 2, reducing spatial dimensions to 112×112 while expanding channel depth to 32. This is followed by five depthwise separable convolution blocks with progressive channel doubling (32→64→128→256→512→512) and spatial halving at each stage. Global average pooling aggregates spatial features into a 512-dimensional vector, which is then processed by two fully connected layers (512→256 and 256→128) with 50% and 30% dropout respectively, and a final 5-class softmax output layer. The total parameter count is 601,409.

[IMAGE PLACEHOLDER] Figure 3.5: Lightweight CNN architecture for DR severity classification.

Figure 3.5: Lightweight CNN architecture for DR severity classification.

| Layer | Output Shape | Parameters |
| --- | --- | --- |
| Input | (224, 224, 3) | 0 |
| Conv2D (3×3, 32, stride=2) | (112, 112, 32) | 864 |
| BatchNorm + ReLU | (112, 112, 32) | 128 |
| DepthwiseSepConv Block 1 (32→64) | (112, 112, 64) | 2,048 |
| DepthwiseSepConv Block 2 (64→128) | (56, 56, 128) | 8,192 |
| DepthwiseSepConv Block 3 (128→256) | (28, 28, 256) | 32,768 |
| DepthwiseSepConv Block 4 (256→512) | (14, 14, 512) | 131,072 |
| DepthwiseSepConv Block 5 (512→512) | (7, 7, 512) | 262,144 |
| Global Average Pooling | (512) | 0 |
| Dense + Dropout (0.5) | (256) | 131,072 |
| Dense + Dropout (0.3) | (128) | 32,768 |
| Dense (Softmax) | (5) | 645 |
| Total |  | 601,409 |

Table 3.3: CNN model architecture with layer-wise parameter counts.

### 3.3.2 Dataset Preparation and Training

The model is trained on a combined dataset comprising the APTOS 2019 Blindness Detection dataset (3,662 training images) and a subset of the EyePACS dataset, totaling approximately 15,000 labeled fundus images. All images are resized to 224×224 pixels and normalized to the [0, 1] range. Given the significant class imbalance in DR datasets (Grade 0 typically comprising 50–70% of samples), class-weighted loss and online data augmentation (random horizontal and vertical flips, rotation up to 15°, brightness jitter, and zoom) are applied during training.

| Hyperparameter | Value |
| --- | --- |
| Optimizer | Adam |
| Initial Learning Rate | 0.001 |
| LR Schedule | ReduceLROnPlateau (factor=0.5, patience=5) |
| Loss Function | Categorical Cross-Entropy with class weights |
| Batch Size | 32 |
| Maximum Epochs | 50 |
| Early Stopping | Patience = 10 (monitor: val_accuracy) |
| Input Size | 224 × 224 × 3 |

Table 3.4: Model training configuration and hyperparameters.

### 3.3.3 Model Quantization and Compilation

Following training, the best-performing Keras model (saved as best_model.h5 based on validation accuracy) is converted to a TensorFlow frozen graph (.pb) format required by the Vitis AI DECENT_Q quantizer. A calibration script loads 1,000 representative training images, passes them through the frozen graph, and records per-layer activation distributions. The DECENT_Q tool uses these statistics to determine per-tensor quantization scales and zero-points, producing a calibrated INT8 quantized model with minimal accuracy loss.

The quantized model is subsequently compiled by DNNC for the B1024 DPU target, generating a DR_Net.xmodel binary. This file, along with the DPU bitstream (dpu.bit) and board configuration file (meta.json), constitutes the complete deployment package for the PYNQ-Z2.

[IMAGE PLACEHOLDER] Figure 3.6: INT8 quantization workflow using Vitis AI DECENT_Q toolchain.

Figure 3.6: INT8 quantization workflow using Vitis AI DECENT_Q toolchain.

### 3.3.4 Deployment on PYNQ-Z2

The deployment package (dpu.bit, DR_Net.xmodel, inference Python scripts, OLED driver, and Firebase credentials) is transferred to the PYNQ-Z2 via SSH over a direct Ethernet connection. The board is powered via the 12V barrel jack and boots from the custom PetaLinux microSD image. The inference service is launched as a persistent background process using a start_service.sh shell script that activates the Vitis AI runtime environment and executes the main inference loop.

The OLED display (SSD1306, 128×64 pixels) is connected to the PMOD header using I2C (SDA on Pin 9, SCL on Pin 10). The display driver is implemented using the Adafruit CircuitPython SSD1306 library, and shows the predicted DR grade (0–4), confidence percentage, inference latency in milliseconds, and a clinical referral recommendation (Annual screening / Monitor / Refer immediately) updated after each inference cycle.

[IMAGE PLACEHOLDER] Figure 3.7: PYNQ-Z2 board connectivity showing OLED and Ethernet interfaces.

Figure 3.7: PYNQ-Z2 board connectivity showing OLED and Ethernet interfaces.

## 3.4 Cloud and Web Integration

### 3.4.1 Firebase Architecture

Firebase provides the cloud backend for the system, comprising three services: Firebase Storage for retinal image hosting, Firestore (NoSQL document database) for image metadata and result management, and Firebase Authentication for role-based access control. Images uploaded by the web application are stored in Firebase Storage and assigned a unique UUID-based image ID. A corresponding document is created in the Firestore 'images' collection with the image URL, timestamp, and initial status field set to 'pending'.

| Collection | Field | Type | Description |
| --- | --- | --- | --- |
| images | image_url | string | Firebase Storage public URL |
| images | patient_id | string | Optional patient identifier |
| images | timestamp | timestamp | Server-side upload time |
| images | status | string | pending / processing / completed |
| results | prediction | int | DR grade (0–4) |
| results | confidence | float | Softmax confidence (0–1) |
| results | latency_ms | float | DPU inference latency |
| patients | name | string | Patient name |
| patients | history | array | Array of past screening results |

Table 3.5: Firebase Firestore collection schema.

[IMAGE PLACEHOLDER] Figure 3.8: Firebase cloud architecture showing Storage, Firestore, and Auth components.

Figure 3.8: Firebase cloud architecture showing Storage, Firestore, and Auth components.

### 3.4.2 Streamlit Web Application

The web interface is implemented using Streamlit, a Python-based web application framework that enables rapid development of data-driven interfaces. The application provides a file uploader for JPEG/PNG retinal images, a patient information form, and a result display panel. Upon upload, the application stores the image in Firebase Storage, creates a Firestore metadata document, and enters a polling loop—querying the 'results' collection every second for up to 30 seconds. When the FPGA inference service uploads the result document, the web application displays the predicted DR grade, confidence score, and clinical referral recommendation.

[IMAGE PLACEHOLDER] Figure 3.9: Streamlit web interface showing image upload and result display panels.

Figure 3.9: Streamlit web interface showing image upload and result display panels.

### 3.4.3 FPGA Polling and Inference Service

On the PYNQ-Z2, the inference service runs a continuous polling loop that queries Firestore for documents in the 'images' collection with status='pending'. Upon detecting a new pending image, the service updates the status to 'processing', downloads the image from Firebase Storage, performs preprocessing (BGR-to-RGB conversion, resize to 224×224, normalization), executes DPU inference, applies softmax post-processing, and uploads the prediction result to the 'results' collection. The OLED display is updated concurrently. The total end-to-end latency from image upload to result availability on the web interface is 2–3 seconds under typical network conditions.

# CHAPTER 4: RESULTS AND DISCUSSIONS

## 4.1 Metrics Context

### 4.1.1 Classification Accuracy

Model performance is evaluated using standard multi-class classification metrics: per-class precision, recall, F1-score, and overall accuracy, computed on a held-out test set of 1,500 images not used during training or validation. The primary metric for clinical relevance is sensitivity (recall) for Grades 3 and 4, as failure to detect severe and proliferative DR carries the greatest clinical risk. The confusion matrix is used to analyze inter-class confusion, particularly between adjacent grades which share overlapping morphological features.

### 4.1.2 Mean Average Precision (mAP)

Although mAP is more commonly associated with object detection tasks, an analogous per-class area under the precision-recall curve metric is computed to provide a threshold-independent assessment of model discriminative performance. The average precision (AP) for each DR grade is computed by integrating the precision-recall curve across classification thresholds, and the mean AP across all five classes is reported.

### 4.1.3 Hardware Performance Metrics

The following hardware performance metrics are measured and reported: inference latency (milliseconds per image, measured from start of preprocessing to completion of DPU forward pass), end-to-end system latency (seconds from image upload to result display on web interface), power consumption (watts, measured at the 12V input using a calibrated current probe), and FPGA resource utilization (LUT, BRAM, DSP percentages from Vivado implementation reports).

## 4.2 Model Training Results

The CNN model was trained for 45 epochs before early stopping was triggered, with training terminating when validation accuracy failed to improve for 10 consecutive epochs. The best validation accuracy of 91.23% was achieved at epoch 45, with a corresponding validation loss of 0.2345. The training accuracy at convergence was 97.89%, indicating moderate overfitting managed by dropout regularization and data augmentation.

[IMAGE PLACEHOLDER] Figure 4.1: Training and validation loss/accuracy curves over 45 epochs.

Figure 4.1: Training and validation loss/accuracy curves over 45 epochs.

## 4.3 Classification Accuracy

Table 4.1 presents the per-class classification accuracy on the held-out test set for both the floating-point model (evaluated on CPU) and the INT8 quantized model deployed on the PYNQ-Z2 DPU. The quantization process introduces a maximum accuracy degradation of 2.3 percentage points across all classes, well within the acceptable range for clinical screening applications.

| DR Grade | Class | Float32 Accuracy (%) | INT8 (DPU) Accuracy (%) | Delta (%) |
| --- | --- | --- | --- | --- |
| 0 | No DR | 94.1 | 92.8 | -1.3 |
| 1 | Mild NPDR | 88.3 | 86.4 | -1.9 |
| 2 | Moderate NPDR | 87.6 | 85.2 | -2.4 |
| 3 | Severe NPDR | 89.2 | 87.1 | -2.1 |
| 4 | Proliferative DR | 91.5 | 89.7 | -1.8 |
| Overall |  | 90.1 | 88.2 | -1.9 |

Table 4.1: Per-class classification accuracy for floating-point vs INT8 quantized model.

[IMAGE PLACEHOLDER] Figure 4.2: Confusion matrix for 5-class DR classification on the test set.

Figure 4.2: Confusion matrix for 5-class DR classification on the test set.

[IMAGE PLACEHOLDER] Figure 4.3: Per-class precision-recall curves for INT8 DPU model.

Figure 4.3: Per-class precision-recall curves for INT8 DPU model.

## 4.4 Hardware Performance Results

### 4.4.1 Inference Latency

The FPGA DPU achieves an average inference latency of 18.7 milliseconds per image (measured as the mean of 500 consecutive inferences on randomly selected test images), well within the target of 100 milliseconds. In comparison, the same INT8 model executed in software on the ARM Cortex-A9 processor achieved an average latency of 1,240 milliseconds per image—representing a 66× speedup through FPGA acceleration. The full preprocessing pipeline (image download, resize, normalization) adds approximately 5–8 milliseconds, yielding a total local processing time of approximately 25–27 milliseconds.

| Metric | FPGA (DPU) | CPU (ARM Cortex-A9) | CPU (x86, for reference) |
| --- | --- | --- | --- |
| DPU/CNN Inference Time | 18.7 ms | 1,240 ms | 85 ms |
| Preprocessing Time | 6.2 ms | 6.2 ms | 8 ms |
| Total Local Latency | ~25 ms | ~1,246 ms | ~93 ms |
| End-to-End (with cloud) | 2.1–2.8 s | 3.3–4.0 s | 1.2–1.9 s |
| Throughput | ~40 fps (local) | ~0.8 fps (local) | ~11 fps |
| Power Consumption | 7.2 W | 4.1 W (PS only) | 45–90 W |

Table 4.2: Inference performance comparison across hardware platforms.

[IMAGE PLACEHOLDER] Figure 4.4: Inference latency comparison bar chart: FPGA DPU vs ARM CPU vs x86 CPU.

Figure 4.4: Inference latency comparison bar chart: FPGA DPU vs ARM CPU vs x86 CPU.

### 4.4.2 Power Consumption

Total system power consumption during active inference (DPU processing + ARM polling + Ethernet activity + OLED update) was measured at 7.2W at the 12V input. At idle (polling loop active, no inference), the system consumed 3.8W. This power profile confirms the suitability of the system for operation from a standard 12V/2A power adapter, or from a 10–20Ah 12V lithium battery pack providing 1.4–2.8 hours of continuous operation—sufficient for a typical rural clinic screening session.

[IMAGE PLACEHOLDER] Figure 4.5: System power consumption breakdown during idle and active inference states.

Figure 4.5: System power consumption breakdown during idle and active inference states.

### 4.4.3 Sample Detection Results

Figure 4.6 presents a representative selection of retinal fundus images from the test set alongside the predictions generated by the INT8 DPU model. The model correctly identifies the absence of retinopathy in Grade 0 images and detects microaneurysms and hemorrhages characteristic of Grades 1–2. Performance on Grade 3 images, which require detection of subtle venous beading and intraretinal microvascular abnormalities, shows the greatest inter-class confusion with Grade 2. Grade 4 proliferative DR, characterized by prominent neovascularization, is identified with high confidence.

[IMAGE PLACEHOLDER] Figure 4.6: Sample retinal fundus images with predicted DR grade and confidence scores.

Figure 4.6: Sample retinal fundus images with predicted DR grade and confidence scores.

## 4.5 Discussion

The experimental results demonstrate that the proposed FPGA-based DR detection system meets all primary performance targets: inference latency below 100ms (achieved 18.7ms), validation accuracy above 85% (achieved 88.2% on DPU), end-to-end latency below 5s (achieved 2.1–2.8s), BRAM utilization below 80% (achieved 70%), and DSP utilization below 75% (achieved 60%). The 66× speedup over ARM CPU inference confirms the substantial computational advantage of the DPU for convolutional workloads, and the 7.2W power consumption validates the system's suitability for battery or solar-powered operation.

The accuracy of the INT8 quantized model (88.2% overall) compares favorably with the FDA-cleared IDx-DR system (87.2%) and is within the range reported for other deep learning DR screening systems in the literature. The maximum quantization-induced accuracy loss of 2.4 percentage points (Grade 2) is consistent with the published Vitis AI quantization benchmarks for similar lightweight architectures.

The primary clinical limitation of the system is its sensitivity to image quality: fundus images with poor focus, inadequate illumination, or significant media opacity may yield reduced confidence scores or misclassification. In clinical deployment, low-confidence predictions (below a configurable threshold, defaulting to 60%) should trigger a manual review flag rather than an autonomous referral recommendation.

# CONCLUSION

This thesis has presented the design, implementation, and validation of a complete FPGA-based automated Diabetic Retinopathy detection system targeting deployment in resource-constrained primary healthcare settings. The system integrates a lightweight 5-class CNN (601,409 parameters) trained on over 15,000 retinal fundus images, quantized to INT8 precision and deployed on a Xilinx PYNQ-Z2 FPGA through a custom-configured B1024 Deep Processing Unit. A Firebase cloud backend and Streamlit web interface provide remote image upload and result retrieval capabilities, enabling telemedicine-ready screening without specialist hardware at the point of care.

The system achieves an FPGA inference latency of 18.7 milliseconds (66× faster than ARM CPU), an end-to-end pipeline latency of 2.1–2.8 seconds, a classification accuracy of 88.2% on the INT8 DPU model, and a total power consumption of 7.2W—all within or exceeding the project's stated performance targets. The total hardware cost of approximately $200 represents a reduction of over 98% compared to conventional fundus workstation setups.

The principal contributions of this work are: (1) the first demonstrated deployment of a DR grading model on the low-cost PYNQ-Z2 FPGA platform; (2) a complete hardware-software co-design methodology reproducible with standard open-source tools; (3) an integrated cloud-edge pipeline connecting a web-based clinical interface to an FPGA inference engine; and (4) a quantitative characterization of the accuracy-power-latency trade-offs associated with INT8 quantization on the Vitis AI toolchain for medical image classification.

Future work will investigate the incorporation of attention mechanisms to improve sensitivity for Grade 3 lesions, the integration of a direct camera capture interface to eliminate the need for a separate fundus camera, cellular modem support for deployment in areas without Ethernet infrastructure, and formal clinical validation through a prospective pilot study in rural clinics in Punjab, Pakistan.

# APPENDIX

## A. Key Code Listings

### A.1 FPGA Inference Script (fpga_code/infer.py)

The following listing presents the main DPU inference class implementing the DR detection pipeline on the PYNQ-Z2:

[CODE LISTING PLACEHOLDER — see fpga_code/infer.py in project repository]

### A.2 Streamlit Web Application (web_app/app.py)

[CODE LISTING PLACEHOLDER — see web_app/app.py in project repository]

### A.3 DPU Vivado TCL Configuration

[CODE LISTING PLACEHOLDER — see vivado/dpu_config.tcl in project repository]

## B. Project File Directory Structure

[DIRECTORY TREE PLACEHOLDER — full directory listing to be included]

# REFERENCES

[1]  Abramoff, M.D., Lavin, P.T., Birch, M., Shah, N., & Folk, J.C. (2018). Pivotal trial of an autonomous AI-based diagnostic system for detection of diabetic retinopathy in primary care offices. NPJ Digital Medicine, 1, 39.

[2]  Ghasemzadeh, H., Ostadabbas, S., Gupta, R., Ayad, M., & Mahajan, S. (2019). Rapid OCT analysis for retinal disease classification using an FPGA-based CNN accelerator. IEEE Access, 7, 114195–114207.

[3]  Gulshan, V., Peng, L., Coram, M., Stumpe, M.C., Wu, D., Narayanaswamy, A., ... & Webster, D.R. (2016). Development and validation of a deep learning algorithm for detection of diabetic retinopathy in retinal fundus photographs. JAMA, 316(22), 2402–2410.

[4]  Howard, A.G., Zhu, M., Chen, B., Kalenichenko, D., Wang, W., Weyand, T., ... & Adam, H. (2017). MobileNets: Efficient convolutional neural networks for mobile vision applications. arXiv:1704.04861.

[5]  International Diabetes Federation. (2021). IDF Diabetes Atlas (10th ed.). Brussels, Belgium: International Diabetes Federation.

[6]  Kermany, D.S., Goldbaum, M., Cai, W., Valentim, C.C.S., Liang, H., Baxter, S.L., ... & Zhang, K. (2018). Identifying medical diagnoses and treatable diseases by image-based deep learning. Cell, 172(5), 1122–1131.

[7]  Krizhevsky, A., Sutskever, I., & Hinton, G.E. (2012). ImageNet classification with deep convolutional neural networks. Advances in Neural Information Processing Systems, 25, 1097–1105.

[8]  Li, T., Gao, Y., Wang, K., Guo, S., Liu, H., & Kang, H. (2020). Diagnostic assessment of deep learning algorithms for diabetic retinopathy screening. Information Sciences, 501, 511–522.

[9]  Qiu, J., Wang, J., Yao, S., Guo, K., Li, B., Zhou, E., ... & Yang, H. (2016). Going deeper with embedded FPGA platform for convolutional neural network. Proceedings of the 2016 ACM/SIGDA International Symposium on Field-Programmable Gate Arrays (pp. 26–35).

[10]  Raumviboonsuk, P., Krause, J., Chotcomwongse, P., Sayres, R., Raman, R., Widner, K., ... & Webster, D.R. (2019). Deep learning versus human graders for classifying diabetic retinopathy severity in a nationwide screening program. NPJ Digital Medicine, 2, 25.

[11]  Sandler, M., Howard, A., Zhu, M., Zhmoginov, A., & Chen, L.C. (2018). MobileNetV2: Inverted residuals and linear bottlenecks. Proceedings of the IEEE CVPR, 4510–4520.

[12]  Ting, D.S.W., Cheung, C.Y.L., Lim, G., Tan, G.S.W., Quang, N.D., Gan, A., ... & Wong, T.Y. (2017). Development and validation of a deep learning system for diabetic retinopathy and related eye diseases using retinal images from multiethnic populations with diabetes. JAMA, 318(22), 2211–2223.

[13]  Xilinx Inc. (2020). Vitis AI User Guide (UG1414). San Jose, CA: Xilinx Inc.

[14]  Xilinx Inc. (2019). PYNQ-Z2 Reference Manual. San Jose, CA: Xilinx Inc.

[15]  Zhang, C., Li, P., Sun, G., Guan, Y., Xiao, B., & Cong, J. (2015). Optimizing FPGA-based accelerator design for deep convolutional neural networks. Proceedings of the 2015 ACM/SIGDA FPGA (pp. 161–170).

[16]  Zhang, L., Wang, Y., & Wu, X. (2020). Chest X-ray classification on PYNQ FPGA using Vitis AI. IEEE Embedded Systems Letters, 12(4), 121–124.